SVM30

# 1. Foreword

The SVM30 is a new board from Sensirion, based on 2 sensor that have been around for a while : SGP30 and SHTC1. The SGP30 provides CO2 and TVOC information, while the SHTC1 is providing temperature and humidity.

The 2 complement each other as humidity information from the SHTC1 can be used to adjust / fine tune the SGP30 output. The board is positioned as a single solution for HVAC systems.

The hardware connection and software function are described in chapter 2 and chapter 3.

While there a number of remarks around the board in chapter 4, the SVM30 is a quality product, as long as it is used in forced air movement, e.g. with a fan.

In chapter 5 the SVM30 a test and compare against CCS811 and SCD30, in Chapter 6 another compare against SCD30, BME280 and BME680

October 2019
Paul van Haastrecht
# 2. Hardware connection

 svm30 pin Raspberry Pi

 1 SCL -------- SCL pin 5 / GPIO 3
 2 GND -------- GND
 3 VCC -------- +5V
 4 SDA -------- SDA pin 3 / GPIO 2

No need for external resistors as pin 3 and 5 have onboard 1k8 pull-up resistors on the Raspberry Pi. However also, and not Documented, but it seems the SVM30 has on-board pull-up resistors (~20K) as the scope picture in Appendix A shows.
# 3. Software
## 3.1 Library Functions (SVM30_lib)
### 3.1.1 void EnableDebugging(bool act)
Enable or disable the printing of library debug messages.
- false : no debug messages
- true : enable debug messages
### 3.1.2 bool begin()
Initialize the communication & start SGP30
return : true on success else false

### 3.1.3 bool probe()
check if SVM30 sensors are available
return : true on success else false

### 3.1.4 const char * GetDriverVersion()
return driver version.
### 3.1.5 bool TriggerSGP30()
Trigger a read on the SGP30

See chapter 3.5 baseline compensation for more info

return : true on success else false

### 3.1.6 bool GetId(uint8_t device, uint16_t *buf)
read ID number from SGP30 or SHTC1
- Parameter device : SGP30 or SHTC1
- Parameter buffer to store ID (min 3 words)
return : true on success else false

### 3.1.7 bool reset(uint8_t device)
Reset SGP30 or SHTC1
Parameter device : SGP30 or SHTC1
return : true on success else false

### 3.1.8 bool GetFeatureSet(char *buf)
read feature set from SGP30
Param buf : buffer to store feature set number (min 2)
return : true on success else false

### 3.1.9 bool MeasureTest()
Perform Measure SELF test on SGP30
return : true on success else false

### 3.1.10 bool GetBaseLine_CO2(uint16_t *baseline)
### 3.1.11 bool GetBaseLine_TVOC(uint16_t *baseline)
get individual baseline from SGP30
return : true on success else false

### 3.1.12 bool GetBaseLines(uint32_t *baseline)
get BOTH baselines (TVOC and CO2) from SGP30
return : true on success else false

### 3.1.13 bool SetBaseLine_CO2(uint16_t baseline)
### 3.1.14 bool SetBaseLine_TVOC(uint16_t baseline)
set individual baseline on SGP30
param baseline (as retrieved with GetBaselines()
source datasheet:
Approximately in the first 15 seconds of operation after start the call will fail unless a previous baseline was restored.
REMARK : Sending baseline values of zero will be treated as an error.
return : true on success else false
### 3.1.15 bool SetBaseLines(uint32_t baseline)
set BOTH baselines (TVOC and CO2) on SGP30
param baseline (as retrieved with GetBaselines()
return : true on success else false

### 3.1.16 bool SetHumidity(float humidity)
set humidity on SGP30
param humidity: humidity value to set
Sending a humidity value of 0x0000 can be used to turn off the humidity compensation. The humidity value used for compensation is set to its default value (11.57g/m3)

return : true on success else false

### 3.1.17 bool GetValues(struct svm_values *v)
read all measurement values from the sensor, perform calculations and store in structure
param v: pointer to structure to store
return : true on success else false

### 3.1.18 void SetTempCelsius(bool act)
Set temperature type
param act :
- true set to Celsius
- false is Fahrenheit

### 3.1.19 close()
Close the library, restore the pins and release memory.
## 3.2 Values available
/* structure to return measurement values */
typedef struct svm_values
{
    uint16_t r_humidity; // SHTC1 (raw data from sensor)
    uint16_t r_temperature; // SHTC1 (raw data from sensor)
    int32_t humidity; // SHTC1 (converted humidity)
    int32_t temperature; // SHTC1 (converted temperature in C)
    float absolute_hum; // calculated absolute humidity in g/m3.1.
    uint16_t CO2eq; // SGP30
    uint16_t TVOC; // SGP30
    uint16_t H2_signal; // SGP30 Raw H2 signal
    uint16_t Ethanol_signal;// SGP30 Raw ethanol signal
    float heat_index; // calculated heat index
    float dew_point; // calculated dew point
};

## 3.3 user level program SVM30
The following options can be used with the user level SVM30 programtriggerSCP30
### 3.3.1 SVM30 settings:
    -c 0x#	set baseline CO2 to ####
    -t 0x#	set baseline TVOC to ####
    -h	continued humidity compensation
    -m	perform a measurement test
### 3.3.2 program control settings:
   -d	display ID-numbers and feature set only
   -l #	number of measurements (0 = endless)
   -w #	wait-time (seconds) between measurements
   -v	include verbose / debug information

### 3.3.3 output formatting
    -D	do not display output in color
    -T	add / remove timestamp
    -H	add / remove humidity & temperature
    -A	add / remove CO2 / TVOC info"
    -B	add / remove baseline info
    -R	add / remove H2 and Ethanol signals
    -E	add / remove Dew point calculation
    -J	add / remove Absolute Humidity calc
    -G	add / remove HeatIndex calc
    -F	Display temperature (Fahrenheit/Celsius)
## 3.4 Compile the program
## 3.4.1 BCM2835 library
Install latest from BCM2835 from : http://www.airspayce.com/mikem/bcm2835/

1. wget http://www.airspayce.com/mikem/bcm2835/bcm2835-1.56.tar.gz
2. tar -zxf bcm2835-1.56.tar.gz	// 1.56 was version number at the time of writing
3. cd bcm2835-1.56
4. ./configure
5. sudo make check
6. sudo make install

In order for this software to run you should NOT enable i2C in raspi-config to load the kernel drivers. It works directly on the hardware, but you’ll have to run program as root.

To create a build with only the SVM30 monitor type:
     make
To create a build with the SVM30 and SDS011 monitor type:
     make BUILD=SDS011 (requires the sds011 sub-directory)

## 3.5 baseline compensation
An interesting aspect of the SGP30 is the ability to self calibrate and adjust the baselines for CO2 and TVOC results
source : Datasheet SVM30:
“ After the “Init_air_quality” command, a “Measure_air_quality” command has to be sent in regular intervals of 1s to ensure proper operation of the dynamic baseline compensation algorithm.”
“ The on-chip baseline compensation algorithm has been optimized for 1HZ sampling rate. The sensor shows best performance when used with this sampling rate.”

Hence you will find in the program a regular triggerSCP30() or GetValues() is called nearly each second.
# 4. Observations
## 4.1 Remarks about SVM30

The SVM30 is actually a combination of 2 “known” units (SHTC1 and SGP30) and as such not really a “new kid on the block”. These are high quality measurement components that have now been combined on the same board.
## 4.2 SVM30 concerns

The claim” Automatic baseline compensation and humidity compensation of MOX gas sensor” is fully correct BUT NOT really what you would expect. (good marketing thou)
- The SGP30 has “ humidity compensation”, but YOU have to read the temperature and humidity from the SHTC1, calculate the absolute humidity and send that in a fixed point format to the SGP30. NOTHING is happening on the board.
The heat dissipation from the on-board components has an impact on the temperature ( < 0.7C) and humidity reading in open air, but in closed cabinet this can have a high impact. The heat dissipation causes an increase of 6 - 7 degrees Celsius on the SGP30, with the transistors in SOT23 package 8 – 9 degrees Celsius warmer than the air (see chapter 4.4)

The board is 5V only VCC. The SGP30 and SHTC1 sensors operate on voltages around 1.8V. Part of the heat dissipation could be because of the needed voltage drop. However the SDA and SCL seem to have on-board pulled-up resistors to 5V. This makes the sensor not usable to 3V3 processor boards without extra hardware.

## 4.3 SGP30 concerns

Receipt of correct baselines or humidity setting is NOT confirmed by the SGP30 sensor. The baselines can be read back, but for humidity it is hopefully accepted. Either way the output results from the SGP30 is not really changing after setting baseline or humidity. The impact is unknown.

Reading the RAW H2 and Ethanol values does NOT really bring any further information and can be totally discarded. Same is true for the self / measurement test. All nice to have

Setting the baseline to a certain value has its wrong or undocumented challenges. It has the same errors and short comings in the datasheet as the SGP30. I wonder why this has not been corrected after 2 years after the latest SGP30 datasheet?? (see chapter 4.5)

The product feature-information decoding of the SGP30 is completely undocumented. If this is planned to provide information about the capabilities of the software in the sensor, why not share what it is ? For this driver I have implemented the info from the datasheet.

## 4.4 Temperature measurement

In general the temperature measurement was correct, in open, none forced moving air. A separate high quality temperature device was used to do compare and even after long term measurement ( >2 hours) and the results were close ( < 0.7 C). Then I connected a external ventilator, an 12V CPU cooler fan, and had it run for hours. The results of both separate temperature devices showed similar results ( < 0.1C difference) . Clearly shows that with air flowing the temperature reading is correct. However when turning off the fan, the temperature increased quickly to 1 degree difference.

One can notice that the temperature around SGP30 is much higher due to heat from the on-board components.(just put your finger on it or hold it against your cheek) Measuring the temperature around the SGP30 with the same high quality temperature device showed a 6 to 7C higher reading, with the transistors in SOT23 even up to 9 degrees warmer.

If this sensor is placed in a cabinet, this could have a high impact and it is extremely important to have to a good shield the SGP30 side of the SVM30 from the SHTC1 side and use a fan to force air flow

## 4.5 Baseline setting

Baseline handling on the SGP30 seems to be half documented and partly documented in error.

1. The order is of reading and sending is TVOC, CO2eq (the SVM30 datasheet is reversing the order up for reading, which is NOT correct)
2. When sending the baseline values, each word needs to be followed with a CRC (not defined in the SVM30 datasheet)
3. Setting a baseline value of zero on CO2 causes CO2 baseline to be the same as TVOC. Setting values of zero should be avoided (not documented)
4. TVOC baseline can be set by sending only 1 word/baseline + CRC. (Not documented)
5. When setting baseline during the first 15 seconds after power-up/reset, update baseline for CO2 first, then TVOC. Else setting will fail.
6. After setting the baseline value during the first 15 seconds after power-up/reset, the baselines will NOT change / updated. Once it is set the baseline stays what it is. (not documented). If set after the initial calibration during the 15 seconds, they will continue to be optimized.

While taken into account the above remarks in chapter 4, the SVM30 a good sensor board.
# 5. Benchmark SVM30 / SCD30 / CCS811

With all the work done, time to put the SVM30 to a benchmark. The other victims are CCS811 and SCD30. All three were connected to a Raspberry PI, for which the software is available

CCS811 : https://github.com/paulvha/ccs811
SCD30 : https://github.com/paulvha/scd30_on_raspberry
SVM30 : https://github.com/paulvha/SVM30_on_raspberry

The test was done with a fan forcing air movement around the 3 sensors. As the CCS811 needs to warm up, the results were taken after 3 hours of working.

## 5.1 Temperature

All 3 provide temperature reading. The CSS811 has an CTN, the SCD30 use an SHT30 and the SVM30 SHTC1. The CSS811 and SCD30 show similar results, although the SCD30 reacts faster. The SVM30 result was consistent lower for the whole time, but it was showing the same results as a high quality temperature meter.
## 5.2 CO2 values

All 3 provide CO2 information. No matter what… the CCS811 was not showing anything else than 400. On a certain moment I have tested the CCS811 (blowing directly on the board) and it moved up for 2 readings, then back to 400. This was before the samples taken below. The SCD30 has consistent higher reading, apart from the samples spikes around 113, where I breathed air in to the fan. Both SVM30 and SCD30 react very fast on changes in the air. Although the SVM30 has humidity compensation is was not turned on.

## 5.3 TVOC values

All CCS811 and SVM30 provide TVOC information. No matter what… CCS811 TVOC continued to be zero.

## 5.4 Compare CO2 and TVOC on SVM30
As one would expect they follow the same pattern, as they use the same sensor, but a different calculation
## 5.5 SVM30 CO2 to baseline

Maybe the baseline helps to get better results for the CO2 reading, the chart below shows that it is not possible to draw conclusion.

## 5.6 SVM30 CO2 to baseline
Comparing the TVOC results and TVOC baseline does not shine a light either

## 5.7 humidity
The SCD30 provide humidity with an SHT30 and the SVM30 with SHTC1. Both quality sensors and sometimes 1% humidity apart.

# 6 Benchmark SCD30 / SVM30 / BME280 / BME680

This benchmark is around humidity, temperature and pressure. All 4 sensors were connected to a Raspberry PI, running at the same time. A computer fan with focused air-stream was used to force air to the 4 sensors. The BME280 and BME680 shared the same I2C line, powering both with 5V.
The sensors have been running for 3 hours, with sampling every minute. The last 1.5 hour samples were used.

## 6.1 Humidity results
Humidity My BME680 RH% were constant much higher. Even moved the BME680 to another place in the air, moved the board around but no difference in the measurement results. On average over the 95 samples, the BME280, SVM30 and SCD30 results are within 3% points. The BME680 is 7% RH higher than SCD30/SVM30.
Average Humidity
## 6.2 Temperature results

Nearly the same as humidity, with the BME680 being higher

Temperature
Looking at the average result over the 1.5 hours, the BME680 is 4% point higher. The BME280, SVM30 and SCD30 are within 2% points.

Average Temperature

## 6.3 Pressure results

Although the SVM30 and SCD30 are not providing pressure information, the BME680 and BME280 do. I have included the information in case you want to add either one into your project:

Pressure

The delta is 1Pa only on average.
# Appendix A: scope signal

This signal was taken WITHOUT external / extra pull-up resistors. The SVM30 must have them on-board, otherwise these signals would not exist or have this quality.